Approximation of temperature at surface

Tsurface=Tdielectric+Tdiffusion

Temperature of surface interface WRT DUE solely to heat diffusion. At x<0 the surface heat is assumed to be a constant 37 since the coaxial cable is conducting heat into the blood circulation/flux at a high rate, the region beneath the catheter is also observed to be heavily spared even though it would have the greatest SAR. So it’s simply assumed that there’s no heat diffusion influence to the surface temperature at x<0

Tdiffusiont=37,  &x<037+βdf(Tcatheter(t)-37),  &x≥0

Where βis an experimentally determined constant, d is distance of catheter axis to surface (m), f is flowing rate of blood(L/min). The temperature at the surface due to diffusion is related to the temperature of the catheter itself, which is given in terms of power, or can be given practically by the thermocouple:

Tcathetar(t)=P1-e-ktmck+T0

Where k is another experimental constant which can be determined by fitting this eqn to the thermocouple measurements since the power drop off is unknown.

Tdielectrict=Ermsσ2ρ

Where ρ = tissue density, σ is the conductivity. E_rms determined via power and wave equations.

Approximation of temperature inside the tissue/gel

Now the heat of the surface can be used as a component of the transient heat equation along with the electric field equations for calculating dielectric heating. The coupled result of these two equations will give the heat at a point inside the gel.

1. The transient heat equation inside the gel

Take x along the catheter from the feed point, z vertical with z = 0 at the gel surface and z < 0 inside the gel (the same axes as the E-field model), and y across the catheter. The gel has no blood perfusion and no metabolic heat, so the Pennes bioheat equation reduces to plain transient conduction with one volumetric source, the dielectric heating:

ρc∂T∂t = k∇2T + Qdx,y,z,t

Here ρ, c and k are the density, specific heat and thermal conductivity of the gel. For a time-harmonic field, the time-averaged power absorbed per unit volume and the SAR are:

Qd = σErms2 = ωε0εg″Erms2  SAR = Qdρ = σErms2ρ

Note on Tdielectric above: the quantity σ|E_rms|²/ρ is the SAR, in W/kg. The square belongs on E_rms, not on σ. The SAR is a heating rate per kilogram, not a temperature. It only becomes a temperature once it is integrated over time. In the first few seconds, before any heat has diffused away, it gives:

Tdielectrict − Tb ≈ SAR·tc = σErms2tρc

After this early period, conduction and the cooled surface pull the temperature below that line. That is why the dielectric term has to enter as the source Q_d in the heat equation, rather than be added to the surface temperature directly.

E_rms comes from the insulated-monopole model, which was computed at the delivered power P_ref (15 W for 50 W set). E_rms scales as √P, so Q_d scales linearly with power. With the generator on for 0 ≤ t ≤ t_on:

Qdx,t = σErms,ref2xPtPref,  Pt = Pset1 − Pdropoff, 0 ≤ t ≤ ton0, t > ton

2. Boundary conditions: where the catheter diffusion term enters

The surface-temperature model above becomes the boundary condition on the top face of the gel. The general form is a convective (Robin) condition. The gel surface exchanges heat with the water flowing over it, and the local temperature of that water is T_diffusion:

−k∂T∂zz=0 = hsTx,y,0,t − Tdiffusionx,t

h_s is the heat-transfer coefficient between the circulating water and the gel surface. As h_s → ∞ the condition becomes T(z = 0) = T_diffusion: the surface is pinned, which is exactly what the surface model assumes. A finite h_s also lets the dielectric heating warm the surface a little, which is what happens in practice.

Restating T_diffusion with the transfer fraction α from the notes (T_blood written as T_b):

Tdiffusionx,t = Tb + αxTcathetert − Tb,  αx = min1, βdf, 0 ≤ x ≤ Lc0, otherwise

This makes two small changes to the version above:

• α is capped at 1, because the surface cannot get hotter than the catheter. The uncapped form β/(d√f) grows without limit as the flow f → 0.

• The catheter ends at its tip, so α drops back to zero beyond x = L_c. Setting L_c very large recovers the original x ≥ 0 form.

The catheter temperature comes from the lumped model in the notes (Catheter Temp 1 and 2). That model is the energy balance m c dT/dt = P_c − k_c(T − T_b). Including the cooling after switch-off, its solution is:

Tcathetert = Tb + Pckc1 − e−t/τc, t ≤ tonTb + Tcatheterton − Tbe−t − ton/τc, t > ton,  τc = mckc

P_c is the power that heats the catheter itself, not the power that reaches the gel. The E-field model assumes that 70 % of the generator power is lost before the antenna. Take P_c to be a fraction γ of that loss, dissipated in the catheter tip: P_c = γ P_set P_dropoff. The notes’ T_0 = 37 °C is T_b here. A thermocouple fit determines only two combinations of these parameters: the steady rise ΔT_c = P_c/k_c and the time constant τ_c = mc/k_c. The GUI can therefore expose those two directly, and m, c, k_c and γ never need to be known separately.

The sides and bottom of the gel block lose heat through the tank walls to the room air:

−k∂T∂n = hwT − Troom

Over a 6-minute run the heat barely reaches these faces, so treating them as adiabatic (h_w = 0) is also acceptable. The initial condition is that the gel has equilibrated with the water:

Tx,y,z,0 = T0 = Tb

3. Splitting the solution into a dielectric part and a diffusion part

The heat equation and all of its boundary conditions are linear in T. This holds as long as k, ρ, c and σ do not change with temperature, which is an assumption: they drift by about 1–2 % per kelvin. Under that assumption the solution splits exactly into two parts, each driven by one heat source. This is the rigorous form of T_surface = T_dielectric + T_diffusion, and it holds at every point in the gel, not just at the surface:

Tx,t = Tb + θdx,t + θsx,t

Dielectric part: heated by Q_d, with the surface water held at T_b:

ρc∂θd∂t = k∇2θd + Qd

−k∂θd∂z = hsθd at z = 0,  θdt = 0 = 0

Diffusion part: no internal source, heated only through the surface by the catheter:

ρc∂θs∂t = k∇2θs

−k∂θs∂z = hsθs − αxTcathetert − Tb at z = 0,  θst = 0 = 0

The split assumes T_0 = T_room = T_b. If they differ, a third part (no source, starting at T_0 − T_b and relaxing toward the walls) is added in the same way. At the surface, with h_s → ∞, θ_d = 0 and θ_s = α(T_catheter − T_b), which recovers the surface equation above. Inside the gel both parts are non-zero. The deep peak comes mainly from θ_d, because the water keeps the surface cool.

4. Analytic check of the diffusion part

Well inside the catheter footprint (0 ≪ x ≪ L_c), with the surface pinned, θ_s depends only on depth. For a sudden surface step of size ΔT, the semi-infinite-solid solution (Carslaw & Jaeger) is:

θz,t = ΔT erfc|z|2at,  a = kρc

The surface actually warms gradually, following the catheter. Duhamel’s theorem handles this by adding up a series of small steps, with Θ(t) = α(T_catheter(t) − T_b):

θsz,t = 0tdΘdτ erfc|z|2at − τ dτ

For gel, a = k/(ρc) ≈ 1.4 × 10⁻⁷ m²/s. The diffusion length √(at) is therefore about 3.6 mm at 92 s and 7.2 mm at 360 s. Heat from the catheter reaches the photographed hot-spot depth (3–4 mm) within the run, which is why the diffusion term matters. This expression is also a quick check on the numerical solver.

5. Solving it (and keeping the GUI fast)

With a realistic E_rms map, θ_d has no closed form, so both parts are solved by finite differences on a grid of spacing Δh. Implicit (backward-Euler) time stepping is stable for any time step, so a step of about 0.5–1 s is set by accuracy alone. An explicit scheme needs:

Δt ≤ Δh22Da,  D = number of dimensions (2 or 3)

For example, Δh = 0.25 mm in 3-D needs Δt ≤ 0.07 s. A 2-D (x, z) slice under the catheter is far faster to solve (one fewer dimension of grid points). It overestimates temperatures, however, because in 2-D the source is effectively infinite in y and heat cannot spread sideways. 3-D is better for final numbers.

Linearity makes the GUI cheap to update. Two stored responses, computed once, cover most slider changes:

• D(x, t): the dielectric part for power switched on at t = 0 and left on, computed at P_ref.

• U(x, t): the diffusion part for a unit step in Θ at t = 0.

Any power, on-time and catheter history then follow by rescaling and superposition:

θdx,t = PdelPrefDx,t − Dx,t − ton,  θsx,t = 0tdΘdτ Ux,t − τ dτ

Here P_del = P_set(1 − P_dropoff), and D(x, t − t_on) = 0 for t < t_on. The sliders then fall into two groups:

• Instant (rescale or convolve only): P_set, P_dropoff, t_on, β, d, f, m, c, k_c, γ, T_b.

• Re-solve (the grid solve must be repeated): the gel’s k, ρ, c and σ, h_s, h_w, L_c, the antenna geometry, and Δh.

6. Starting values for the GUI (first guesses)

None of these values is fitted yet. Each is a guess from the setup table, the photographed run (50 W CW, 6 min) or the V4 model defaults, and each is meant to be replaced as data come in.

| Symbol | Meaning | Guess | Basis |
| --- | --- | --- | --- |
| Power and timing |  |  |  |
| Pset | Generator setting | 50 W | Photographed run |
| Pdropoff | Fraction lost before the antenna | 0.7 | Setup table; tunable |
| ton | Power-on time | 360 s | 6 min run |
| Pref | Power at which E_rms was computed | 15 W | insulated_monopole_Erms.m |
| Gel |  |  |  |
| ρ | Density | 1000 kg/m³ | 1 % Phytagel, >98 % water (V4) |
| c | Specific heat | 4180 J/(kg K) | Water (V4) |
| k | Thermal conductivity | 0.60 W/(m K) | Water (V4) |
| σ | Effective conductivity at 2.45 GHz | 1.84 S/m | ωε₀ε″ with ε″ = 13.5 (table quotes 1.7 S/m) |
| Tb, T0 | Water / initial gel temperature | 37 °C | Setup table (photographed day was 25 °C) |
| Boundaries |  |  |  |
| hs | Water-to-gel surface coefficient | 1400 W/(m² K) | Laminar flat plate at 5 L/min (V4); ∞ pins the surface |
| hw | Tank wall to room | 8 W/(m² K) | Acrylic wall + free convection (V4); 0 = adiabatic |
| Troom | Room air | 25 °C | Lab |
| Surface transfer (diffusion term) |  |  |  |
| d | Catheter axis to gel surface | 2 mm | Setup table (photo shows 1.1 mm) |
| f | Water flow rate | 5 L/min | Setup table, circulation |
| β | Transfer constant | 9 × 10⁻⁴ m·(L/min)^½ | Chosen so α = 0.2: the photos show a cool top layer, so the surface is only mildly warmed |
| Lc | Heated catheter length | 12 mm | Photographed element length |
| Catheter lumped model |  |  |  |
| m | Heated catheter-tip mass | 2 g | Tip + water in the lumen; guess |
| c | Catheter specific heat | 1500 J/(kg K) | Polymer/copper/water mix; guess |
| kc | Catheter-to-water loss coefficient | 0.1 W/K | h ≈ 900 W/(m² K) over a 3 mm × 12 mm tip |
| γ | Share of lost power heating the tip | 0.043 | Gives P_c = 1.5 W, so ΔT_c = 15 K (tip ≈ 52 °C) |
| τc | Catheter time constant | 30 s | = mc/k_c; replace with the thermocouple fit |
| Numerics |  |  |  |
| Δh, Δt | Grid spacing, time step | 0.25 mm, 0.5 s | Implicit scheme |

Sanity check on these guesses. At 3 mm depth the E-field model gives E_rms ≈ 1.2 kV/m, so Q_d ≈ 2.7 MW/m³. The early-time (adiabatic) rise is then about 0.65 K/s. The catheter guesses add a surface drive of αΔT_c ≈ 3 K. The dielectric term should therefore dominate the deep hot spot, and the catheter term should mainly warm the top few millimetres beyond the feed.

Sources. Pennes, J. Appl. Physiol. 1, 93–122 (1948), for the bioheat equation. Incropera & DeWitt, Fundamentals of Heat and Mass Transfer, for conduction, convective boundary conditions and the flat-plate h. Carslaw & Jaeger, Conduction of Heat in Solids, 2nd ed. (1959), for the erfc solution and Duhamel’s theorem. Property values follow Gel_model_V4/model_defaults.m.